St. Martin’s Marsh and Big Bend Seagrasses Aquatic Preserves (SMMAP & BBSAP)

Water Quality Metadata Report

January - December 2009
Latest Update: 03/29/2010

I. Data set and Research Descriptors

1) Principal investigator & contact persons:

Principal Investigator:

Jonathan Brucker, B.S., Environmental Specialist I

3266 N. Sailboat Ave

Crystal River, FL 34428

(352) 563-0450

jonathan.brucker@dep.state.fl.us

Other Contact Persons:

Melissa Charbonneau, Environmental Specialist III, Manager

E-mail:  melissa.charbonneau@dep.state.fl.us  Phone: (352) 563-0450

Timothy Jones, Environmental Specialist II

Email:  timothy.w.jones@dep.state.fl.us  Phone: (352) 563-0450

2) Entry verification:

Data from each deployment are uploaded from the YSI data sonde to a PC, and graphs are produced using EcoWatch software Version 3.18.  These graphs are examined for suspect data that, for example, might derive from probe failure.  Notes are made of any abnormal data, and sensors are reconditioned as necessary.  The files are subsequently exported from EcoWatch in a comma-delimited format (.csv) and opened in Microsoft Excel for processing.  Two copies of this file are saved in the .csv format; one file is used to archive the raw data, while the second file is edited to adjust all records to Eastern Standard Time (as needed), to adjust the time as needed to read exactly at the top or bottom of the hour (e.g., 00:00:00), to remove headers and footers, and to delete data at the beginning and end of each deployment when the sondes were out of the water.  These latter data are identified by field notes that document the times in and out of the water, together with depth and conductivity values near zero.

Using the edited .csv files, monthly data are compiled and processed by various macros created by Aquatic Preserve staff and the National Estuarine Research Reserve’s Central Data Management Office (CDMO).  These macros are designed to: 1) check data files for missing data points, 2) fill all cells that do not contain data with periods (.), 3) convert the data columns to appropriate formats for time, date and numeric variables, 4) find all data that are out of acceptable range for the sensors, 5) save the files as Excel worksheets (*.xls), and 6) generate single parameter monthly graphs.

Anomalous data are evaluated to determine if the suspect data should be deleted.  Data are flagged if the values are: 1) outside the range expected for the site, or 2) outside the range of measurements established for the sensors (see Table 1).  Data outside the "normal" range of water quality parameters for a particular site were investigated for validity based on field observations, QC checks, PC6000 printouts, and instrument diagnostics.  Data are deleted if the anomalies are attributed to sensor malfunction.  In addition to observations of any physical damage (e.g., a torn DO probe membrane), sensor malfunctions are detected if the voltage reading of the probe is outside the range established for the sensor or the sensor will not calibrate.  Jonathan Brucker is responsible for these tasks.

3) Research Objectives:

Water quality monitoring using YSI dataloggers began within the SMMAP and BBSAP in early 2004.  Currently, four stations exist where YSI 6600 EDS data sondes have been deployed to continuously monitor selected abiotic parameters at 15-min intervals.  In addition to these four sites, four stations equipped with YSI 600 OMS data sondes are also monitoring selected abiotic parameters at 15-min intervals.  Historically, all data was recorded at 30-min intervals.  The positions of these stations allow for comparisons between relatively pristine versus more urbanized drainage basins as well as higher versus lower salinity regions of the estuary (see “Station Descriptions” under “Site Location and Character”).  The objective of this effort is to quantify the spatial/temporal variability and trends, both seasonally and as a function of tidal forcing, of selected abiotic parameters (e.g. establish baseline data) within the Preserve.

4) Research methods:

YSI 600 OMS data sondes have been continuously operated (data collection interval: 30 minutes) at the Crystal River site since January 2004, at the Bennett Creek, King’s Bay, and Homosassa River monitoring stations since February 2004, and at the Withlacoochee monitoring station since January 2004.  YSI 6600 EDS data sondes were operated at Cat Island and Lone Cabbage Key beginning in March 2004, and at Seahorse Key and Gomez Rocks beginning in April 2004.  These models incorporate a specially designed wiper apparatus attached to the turbidity probe that reduces the oxygen and pH sensor fouling and thereby improves the quality of data collected.  At each site, the sonde is contained within a 10 cm (inside diameter) housing pipe mounted vertically on a piling.  To facilitate water flow across the sensors, several 2 cm diameter holes were drilled into the submerged portion of the pipe.  Hole density is greatest near the base where the sonde sensors are located.  In early 2005, the Cat Island and Lone Cabbage Key sites were removed.  During most of 2005 and early 2006, data was sporadically collected at all sites due to a lack of staff.

As of March 2006, all YSI 600 OMS sites were operational.  In July 2006, the Gomez Rocks site was removed, and in August 2006, the Seahorse Key station became operational.  In October 2006, the YSI 600 OMS was replaced with an YSI 6600 EDS sonde at the King’s Bay station.  In March 2007, a 6600 EDS station was established in Dekle Beach.  In early 2009, the data collection interval was changed to 15 minutes at all locations.  In March 2009, an additional 6600 EDS station was installed at the mouth of the Suwannee River.  At this time, additional 6600 sites are being determined and planned.

Data recorded on the 600’s are downloaded on-site at approximately two week intervals.  The sondes are inspected for damage and fouling, and if no maintenance is required, the sondes are cleaned, reactivated, and returned to the water.  The 600’s are retrieved and returned to the laboratory for cleaning, calibration, and any additional maintenance throughout the year.

Sonde exchanges at the 6600 EDS sites are made at approximately two week intervals.  At the end of a sampling period, sondes are returned to the laboratory where post-deployment readings and reconditioning take place in accordance with the methods outlined in the YSI Operating and Service Manual.  The EDS turbidity wiper brush is removed and replaced with a clean wiper to avoid contamination of standards during post-deployment procedures.  After a superficial rinse of the sonde in tap water, post deployment readings are recorded for pH (Fisher 7.00 buffer solution) and specific conductivity (Exaxol 50.00 mS/cm standard); a post-deployment turbidity reading in 0.0 NTU standard (DI water) is recorded after a more thorough rinse of the turbidity sensor.  The results of these post-deployment readings are used to evaluate the validity of data (See Tables 3, 4, and 5).

5)  Site location and character:

The St. Martins Marsh Aquatic Preserve was established on October 21, 1969.  The St. Martins Marsh Aquatic Preserve covers open water areas from the Crystal River to the Homosassa River in coastal Citrus County.  It is composed of open water, several inlet bays, tidal rivers and creeks, salt marsh, and adjoins upland hammock islands.  Nutrient exchange between the marshes and the Gulf of Mexico makes the salt marsh a significant area of primary production and a nursery ground for commercial and recreational fish species.  St. Martins Marsh Aquatic Preserve’s freshwater tributaries includes two 1st magnitude spring fed rivers; the Homosassa River to the south and Crystal River to the north.  Spring discharge does not fluctuate dramatically from season to season allowing a constant flow of freshwater into St. Martins Marsh’s productive and well-balanced estuary.  The area’s vast coastal salt marshes, mud flats, oyster bars, mangrove islands, and seagrass beds are the southern terminus for migratory waterfowl of the Atlantic and Mississippi flyways.  St. Martins Marsh provides stopover and wintering areas for many migratory species.  The Springs Coast is characterized by unique limestone outcroppings and exposed karstic features.  Habitats associated with these areas are hard-bottom sponge and coral communities and Sargassum meadows.

The Big Bend Seagrasses Aquatic Preserve was established in 1985.  Its boundaries extend from the Withlacoochee River to the St. Mark’s River and out nine nautical miles.  The preserve boundary encompasses all tidal lands, islands, seagrass beds, shallow banks, and submerged bottoms from the mean high water line extending nine miles into the Gulf of Mexico. Landward, it includes all natural waterways tidally connected to the preserve to the extent of state jurisdiction.  Spanning over 945,000 acres, the Big Bend Seagrasses Aquatic Preserve is the largest aquatic preserve and one of the most pristine places in Florida.  The Big Bend Seagrasses Aquatic Preserve consists mainly of a large, remote, and undeveloped expanse of submerged seagrasses and nearshore marshlands located along some 150 miles of the northeast coast of the Gulf of Mexico where the Florida peninsula joins the panhandle.  Numerous estuaries, which nurture a diverse flora and fauna, are formed at the confluence of the many rivers and streams that flow into the preserve.  Open waters and submerged bay bottoms of these estuaries provide habitat to a wide variety of sea and shore birds, including a large nesting colony of frigate birds on Seahorse Key.  This region supports a very important commercial shellfish industry including Cedar Key clams, scallops, oysters, pink shrimp, and blue crab; this area of Florida is a popular destination for the recreational scallop season.  The Suwannee River region supports Essential Fish Habitat (EFH) and the most viable population of the threatened gulf sturgeon.  Big Bend’s vast seagrass beds with mud and sand substrates are important marine habitats to this species.

Station Descriptions:

Table 1: Station Descriptions

| Site Name | Identification Code | Latitude and Longitude | Site Description |
| --- | --- | --- | --- |
| Bennett Creek | BC | N29 01.078 W82 43.621 | Sand/rock bottom, near residential development |
| Crystal River | CR | N28 55.502 W82 41.227 | Sand/mud bottom, adjacent to oyster bars, near mouth of river |
| Dekle Beach | DB | N29 48.836 W83 37.735 | Sand/rock bottom, located on Channel Marker #1 |
| Homosassa River | HS | N28 46.224 W82 41.783 | Sand/mud/rock bottom, near mouth of river |
| King’s Bay | KB | N28 53.105 W82 35.827 | Sand/rock bottom, near residential developments, high tourism activity |
| Seahorse Key | SK | N29 06.109 W83 04.588 | Sand/seagrass bottom, NE corner of Seahorse Key |
| Suwannee River | SW | N29 17.269 W83 09.965 | Mud bottom, at the mouth of the West Pass of the Suwannee River |
| Withlacoochee | WT | N29 00.063 W82 45.422 | Sand/rock bottom, near mouth of river |

6) Data collection period:

Individual sonde deployment and retrieval dates and times for 2009 data are as follows:

BEGAN			ENDED

Bennett Creek Site

12/19/08, 10:30:00		01/05/09, 10:30:00

01/05/09, 10:45:00		01/27/08, 14:00:00

01/27/09, 14:30:00		02/16/09, 10:00:00

02/16/09, 10:15:00		03/06/09, 12:30:00

03/06/09, 12:45:00		03/19/09, 09:30:00

03/19/09, 09:45:00		04/08/09, 09:15:00

04/08/09, 09:30:00		04/22/09, 09:30:00

04/22/09, 09:45:00		05/29/09, 08:30:00

05/29/09, 08:45:00		06/12/09, 10:45:00

06/12/09, 11:00:00		07/08/09, 09:00:00

07/08/09, 09:15:00		07/29/09, 10:15:00

07/29/09, 10:30:00		08/19/09, 10:45:00

08/19/09, 11:00:00		09/10/09, 08:30:00

09/10/09, 08:45:00		09/28/09, 09:15:00

09/28/09, 09:30:00		10/16/09, 12:45:00

10/16/09, 13:00:00		11/02/09, 13:30:00

11/02/09, 13:45:00		11/19/09, 13:15:00

11/19/09, 13:30:00		12/09/09, 09:45:00

12/09/09, 10:00:00		01/14/10, 11:30:00

Crystal River Site

12/31/08, 10:00:00		01/22/09, 13:30:00

01/22/09, 14:00:00		02/18/09, 10:45:00

02/18/09, 11:00:00		03/11/09, 12:00:00

03/11/09, 12:15:00		03/26/09, 08:15:00

03/26/09, 08:30:00		04/16/09, 08:15:00

04/16/09, 08:30:00		05/26/09, 13:00:00

05/26/09, 13:15:00		06/17/09, 08:30:00

06/17/09, 08:45:00		07/15/09, 10:15:00

07/15/09, 10:30:00		08/04/09, 08:30:00

08/04/09, 08:45:00		08/18/09, 09:45:00

08/18/09, 10:00:00		09/17/09, 12:00:00

09/17/09, 12:15:00		09/30/09, 12:30:00

09/30/09, 12:45:00		10/22/09, 07:45:00

10/22/09, 08:00:00		11/12/09, 10:15:00

11/12/09, 10:30:00		12/10/09, 11:00:00

12/10/09, 11:15:00		01/20/10, 14:00:00

Dekle Beach Site

12/10/08, 11:30:00		01/08/09, 11:30:00

01/08/09, 11:45:00		02/26/09, 12:00:00

02/26/09, 12:15:00		03/31/09, 10:00:00

03/31/09, 10:15:00		05/11/09, 10:30:00

05/11/09, 11:00:00		06/01/09, 14:30:00

06/01/09, 14:45:00		07/14/09, 12:00:00

07/14/09, 12:15:00		07/30/09, 11:30:00

07/30/09, 11:45:00		08/28/09, 09:45:00

08/28/09, 10:00:00		10/05/09, 10:00:00

10/05/09, 10:30:00		10/28/09, 10:30:00*

10/28/09, 10:45:00		11/13/09, 12:30:00

11/13/09, 12:45:00		12/11/09, 12:45:00

12/11/09, 13:00:00		01/12/10, 13:00:00

Homosassa River Site

12/12/08, 10:30:00		01/07/09, 14:00:00

01/07/09, 14:30:00		02/06/09, 14:00:00

02/06/09, 14:15:00		03/04/09, 14:15:00

03/04/09, 14:45:00		03/26/09, 13:30:00*

03/26/09, 15:00:00*		04/15/09, 12:30:00*

04/15/09, 13:00:00		05/26/09, 09:00:00

05/26/09, 09:15:00		06/15/09, 09:45:00

06/15/09, 10:00:00		07/07/09, 09:30:00

07/07/09, 09:45:00		08/05/09, 11:45:00

08/05/09, 12:00:00		08/26/09, 08:15:00

08/26/09, 08:30:00		09/21/09, 13:30:00

09/21/09, 13:45:00		10/09/09, 09:00:00

10/09/09, 09:15:00		10/26/09, 12:30:00

10/26/09, 12:45:00		11/12/09, 13:15:00

11/12/09, 13:30:00		11/30/09, 13:45:00

11/30/09, 14:00:00		12/16/09, 11:45:00

12/16/09, 12:00:00		01/26/09, 10:15:00

King’s Bay Site

12/31/08, 09:30:00		01/22/09, 14:00:00

01/22/09, 14:30:00		02/18/09, 11:15:00

02/18/09, 11:30:00		03/11/09, 12:30:00

03/11/09, 12:45:00		03/26/09, 07:45:00

03/26/09, 08:00:00		04/16/09, 08:45:00

04/16/09, 09:00:00		05/27/09, 13:15:00

05/27/09, 13:45:00		06/17/09, 09:00:00*

06/17/09, 09:15:00		07/15/09, 10:45:00

07/15/09, 11:00:00		08/04/09, 08:00:00

08/04/09, 08:15:00		08/18/09, 10:15:00

08/18/09, 10:30:00		09/17/09, 12:30:00

09/17/09, 12:45:00		09/30/09, 13:00:00

09/30/09, 13:15:00		10/22/09, 08:15:00

10/22/09, 08:30:00		11/12/09, 10:45:00

11/12/09, 11:00:00		12/10/09, 11:30:00

12/10/09, 11:45:00		01/18/10, 11:30:00

Seahorse Key Site

12/30/08, 10:30:00		01/26/09, 09:30:00

01/26/09, 09:45:00		03/05/09, 10:00:00

03/05/09, 10:15:00		04/01/09, 09:00:00

04/01/09, 09:15:00		04/28/09, 09:30:00

04/28/09, 09:45:00		05/28/09, 11:15:00

05/28/09, 11:30:00		06/19/09, 10:45:00*

06/19/09, 11:00:00		07/17/09, 09:30:00

07/17/09, 09:45:00		08/05/09, 08:30:00

08/15/09, 08:45:00		08/25/09, N/A*

09/08/09, 09:30:00		10/18/09, 09:00:00

10/08/09, 09:15:00		10/29/09, 09:15:00

10/29/09, 09:30:00		11/16/09, 13:00:00

11/16/09, 13:15:00		12/15/09, 12:30:00

12/15/09, 12:45:00		01/25/10, 10:45:00

Suwannee River Site

03/12/09, 11:45:00		03/31/09, 13:00:00

03/31/09, 13:15:00		05/11/09, 13:00:00

05/11/09, 13:15:00		06/01/09, 12:00:00

06/01/09, 12:15:00		07/14/09, 09:45:00

07/14/09, 10:00:00		07/30/09, 09:30:00

07/30/09, 09:45:00		08/31/09, 09:30:00*

08/31/09, 09:45:00		10/01/09, 12:15:00

10/01/09, 12:30:00		10/28/09, 12:30:00

10/28/09, 12:45:00		11/13/09, 10:30:00

11/13/09, 10:45:00		12/11/09, 10:15:00

12/11/09, 10:30:00		1/12/10, 10:30:00

Withlacoochee Site

01/27/09, 14:00:00		02/16/09, 09:45:00

02/16/09, 10:00:00		03/06/09, 12:00:00

03/06/09, 12:15:00		03/19/09, 09:00:00

03/19/09, 09:15:00		04/08/09, 08:45:00

04/08/09, 09:00:00		04/22/09, 09:00:00

04/22/09, 09:15:00		05/29/09, 07:45:00

05/29/09, 08:00:00		06/12/09, 10:00:00

06/12/09, 10:15:00		07/08/09, 08:45:00

07/08/09, 09:00:00		07/29/09, 09:30:00

07/29/09, 10:00:00		08/19/09, 10:15:00

08/19/09, 10:30:00		09/10/09, 08:15:00

09/10/09, 08:30:00		09/28/09, 09:00:00

09/28/09, 09:15:00		10/16/09, 13:00:00

10/16/09, 13:15:00		11/02/09, 13:15:00

11/02/09, 13:30:00		11/19/09, 13:00:00

11/19/09, 13:15:00		12/09/09, 09:30:00

12/09/09, 09:45:00		01/14/10, 11:15:00

* indicates loss of data due to battery failure, out of water for cleaning, and/or other internal problems that occurred during deployment.

7) Distribution:

The Principle Investigator (PI) retains the right to be fully credited for having collected and processed the data.  Following academic courtesy standards, the PI and Aquatic Preserve (AP) site where the data were collected will be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  Water quality data and metadata can be obtained from the PI (see section 1).

8) Associated researchers and projects:

The SMMAP and BBSAP have formed partnerships with other agencies and organizations actively involved in resource protection in the preserves’ watershed.  Cooperating managers of lands within the AP’s include the:  1) National Park Service (NPS), 2) Suwannee River Water Management District (SRWMD), 3) Southwest Florida Water Management District (SWFWMD), 4) Department of Environmental Protection (DEP) Division of Recreation and Parks/Florida Park Service (FPS), 5) DEP Aquatic Preserve Program, 6) Florida Fish and Wildlife Conservation Commission (FWC), 7) Florida Division of Forestry (DOF), and 8) Citrus, Taylor, Jefferson, Dixie, Levy, and Wakulla Counties.

Other water quality research and monitoring initiatives within the Aquatic Preserves include:  Project COAST in conjunction with the laboratory of Dr. Thomas Frazer at the University of Florida [This work includes water quality, light penetration (PAR), and nutrient analyses.

II. Physical Structure Descriptors

9) Sensor Specifications:

### Table 2.  YSI 6600 EDS data sonde

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Thermistor

Model #: 6560

Range: -5 to 45 °C

Accuracy: +/-0.15 °C

Resolution: 0.01 °C

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: 4-electrode cell with autoranging

Model #: 6560

Range: 0 to 100 mS/cm

Accuracy: +/-0.5% of reading + 0.001 mS/cm

Resolution: 0.001 mS/cm to 0.1 mS/cm (range dependent)

Parameter: Salinity

Units: parts per thousand (ppt)

Sensor Type: Calculated from conductivity and temperature

Range: 0 to 70 ppt

Accuracy: +/- 1.0% of reading or 0.1 ppt, whichever is greater

Resolution: 0.01 ppt

Parameter: Dissolved Oxygen % saturation

Units: percent air saturation (%)

Sensor Type: Rapid Pulse – Clark type, polarographic

Model #: 6562

Range: 0 to 500 % air saturation

Accuracy: 0-200 % air saturation, +/- 2 % of the reading or 2 % air saturation, whichever is greater; 200-500 % air saturation, +/- 6 % of the reading

Resolution: 0.1 % air saturation

Parameter: Dissolved Oxygen mg/L (Calculated from % air saturation, temperature and salinity)

Units: milligrams per Liter (mg/L)

Sensor Type: Rapid Pulse – Clark type, polarographic

Model #: 6562

Range: 0 to 50 mg/L

Accuracy: 0 to 20 mg/L, +/- 2 % of the reading or 0.2 mg/L, whichever is greater; 20 to 50 mg/L, +/- 6 % of the reading

Resolution: 0.01 mg/L

Parameter: Non-Vented Level – Shallow (Depth)

Units: feet or meters (ft or m)

Sensor Type: Stainless steel strain gauge

Range: 0 to 30 ft (9.1 m)

Accuracy: +/- 0.06 ft (0.018 m)

Resolution: 0.001 ft (0.001 m)

Parameter: pH (specify whether EDS probe or not)

Units: units

Sensor Type: Glass combination electrode

Model #: 6561

Range: 0 to 14 units

Accuracy: +/- 0.2 units

Resolution: 0.01 units

Parameter: Turbidity

Units: nephelometric turbidity units (NTU)

Sensor Type: Optical, 90 ° scatter, with mechanical cleaning

Model #: 6136

Range: 0 to 1000 NTU

Accuracy: +/- 5 % reading or 2 NTU (whichever is greater)

Resolution: 0.1 NTU

Dissolved Oxygen Qualifier:  The reliability of the dissolved oxygen (DO) data after 96 hours post-deployment for non-EDS (Extended Deployment System) data sondes may be problematic due to fouling which forms on the DO probe membrane during some deployments (Wenner et al. 2001).  Many Aquatic Preserves have upgraded to YSI 6600 EDS data sondes, which increase DO accuracy and longevity by reducing the environmental effects of fouling.  The user is therefore advised to consult the metadata and to exercise caution when utilizing the DO data beyond the initial 96-hour time period.  However, this potential drift is not always problematic for some uses of the data (e.g., periodicity analysis). It should be noted that the amount of fouling is site specific and that not all data are affected.  The Principal Investigator at should be contacted concerning the reliability of the DO data because of the site and seasonal variation in the fouling of the DO sensor.

Depth Qualifier:  The water quality monitoring program utilizes YSI data sondes that can be equipped with either depth or water level sensors.  Both sensors measure water depth, but by convention, level sensors refer to atmospherically vented measurements and depth refers to non-vented measurements.  Standard calibration protocols for the non-vented sensor use the atmosphere pressure at the time of calibration.  Therefore, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.03 cm for every 1millibar change in atmospheric pressure.  This error is eliminated for level sensors because they are vented to the atmosphere throughout the deployment time interval.  If proper atmospheric pressure data is available, non-vented sensor depth measurements can be corrected for deployments between calibrations. Readings for both vented and non-vented sensors are automatically compensated for water density changes due to variations in temperature and salinity.  The Principal Investigator should be contacted in order to obtain information regarding atmospheric pressure data availability. All data sondes used at all 6600 sites are  non-vented models.

10) Coded variable definitions:

BC = Bennett Creek; CR = Crystal River; DB = Dekle Beach; HS = Homosassa River; KB = King’s Bay; SK = Seahorse Key; SW = Suwannee; WT = Withlacoochee.

11) Anomalous/Suspect, Deleted, and Missing data:

Note #1: Slight shifts in data are sometimes correlated with sonde exchanges.  These shifts are most noticeable in pH, specific conductivity, salinity, DO% and DO conc, and may be related to sensor drift (e.g., due to fouling) and/or calibration/performance differences between sondes.

Note #2: Unless specifically stated in Section 11, turbidity “outliers” (i.e., values that are negative or greater than 1000 NTU) were not deleted from the monthly records.  Negative turbidity values occur throughout the year at all four 6600 EDS sites.  Some of these negative values are within the accuracy range of the sensor (+/- 2.0 NTU) and, therefore, were not removed from the dataset.

Note #3:  All times in data files at all sites had to be adjusted post-deployment; more specifically, times were altered such that the readings occurred on the hour, quarter hours, and half hour.  It has been determined that this clock error is a software issue and has been resolved.

January 1-31, 2009

BC

- Data collection interval was changed to 15 minutes as of 01/05/09 10:45:00; therefore, data is missing intermittently between 01/01/09 00:00:00 and 01/05/09 10:45:00 in the monthly files.

- Missing data 01/25/09 06:45:00, 07:30:00, 09:30:00, 10:00:00, 10:30:00, 11:00:00, 12:30:00 - 12:45:00, 13:15:00, 13:45:00 – 14:00:00, 15:00:00 – 15:15:00, 15:45:00 – 16:00:00, 16:30:00 – 17:00:00, 17:30:00 – 19:00:00, 19:30:00 – 20:30:00, 01/25/09 21:00:00 – 01/27/09 14:00:00, 01/28/09 03:15:00, 04:15:00 – 05:00:00, 05:30:00, 06:15:00, 06:45:00 – 07:15:00, 07:45:00, 08:45:00 – 09:30:00, 10:00:00, 10:30:00 – 11:00:00, 11:30:00 – 12:45:00, 13:15:00, 13:45:00, 14:15:00, 14:45:00 – 16:45:00, 17:15:00, 18:30:00 – 20:45:00, 01/28/09 21:15:00 – 01/29/09 03:15:00, 01/29/09 03:45:00 – 05:00:00, 05:30:00 – 12:15:00, 12:45:00, and 13:15:00 – 01/31/09 23:45:00; battery failure.

CR

- Data collection interval was changed to 15 minute as of 01/22/09 13:45:00; therefore data is missing intermittently between 01/01/09 00:00:00 and 01/22/09 13:45:00 in the monthly files.

DB

- Data collection interval was changed to 15 minutes as of 01/08/09 11:15:00; therefore, data is missing intermittently between 01/01/09 00:00:00 and 01/08/09 11:15:00 in the monthly files.

- Suspect intermittent negative turbidity data between 01/01/09 06:30:00 and 01/06/09 18:00:00; the post deployment value was 1.0 NTU in 0.0 NTU water, and therefore the negative turbidity data may be attributed to a calibration offset.

- Suspect turbidity data 01/07/09 06:00:00, 01/08/09 05:30:00, and 01/12/09 09:15:00; values out of range, reason unknown.

HS

- Data collection interval was changed to 15 minutes as of 01/07/09 14:15:00; therefore, data is missing intermittently between 01/01/09 00:00:00 and 01/07/09 14:15:00 in the monthly files.

KB

- Data collection interval was changed to 15 minutes as of 01/22/09 14:15:00; therefore, data is missing intermittently between 01/01/09 00:00:00 and 01/22/09 14:15:00 in the monthly files.

SK

- Data collection interval was changed to 15 minutes as of 01/26/09 09:15:00; therefore, data is missing intermittently between 01/01/09 00:00:00 and 01/26/09 09:15:00 in the monthly files.

- Deleted all data 01/01/09 08:30:00 – 11:30:00, 01/08/09 04:00:00 – 07:30:00, 01/09/09 04:30:00 – 09:00:00, 01/10/09 06:00:00 – 09:00:00, 01/11/09 07:00:00 – 09:30:00, 01/12/09 07:00:00 – 11:00:00, 01/13/09 08:30:00 – 10:30:00, 01/14/09 08:30:00 – 11:30:00, 01/15/09 10:00:00 – 11:30:00 and 22:30:00 – 01/16/09 01:30:00, 01/16/09 08:30:00 – 12:30:00 and 22:00:00 – 01/17/09 02:00:00, 01/21/09 02:00:00 – 08:30:00, 01/22/09 04:30:00 – 07:00:00, 01/23/09 05:30:00 – 07:30:00, 01/24/09 06:00:00 – 08:00:00, 01/25/09 06:30:00 – 08:30:00, 01/26/09 07:00:00 – 09:00:00, 01/27/09 07:00:00 – 09:30:00, 01/30/09 09:30:00 – 10:15:00 and 21:15:00 – 01/31/09 00:15:00; data suggests sonde was out of the water.

SW

- Missing data 01/01/09 00:00:00 – 01/31/09 23:45:00; sonde not deployed.

WT

- Data collection interval was changed to 15 minutes as of 01/27/09 13:45:00; therefore, data is missing intermittently between 01/01/09 00:00:00 and 01/27/09 13:45:00 in the monthly files.

February 1-28, 2009

BC

- Missing data 02/01/09 00:00:00 – 02/16/09 10:00:00; battery failure.

CR

No data anomalies.

DB

- Shift in D.O. 02/26/09 12:15:00 associated with calibration offset (Post deployment readings were 134.7 and 134.6).

- Suspect turbidity data 02/14/09 12:15:00; value out of range, reason unknown.

HS

No data anomalies.

KB

No data anomalies.

SK

- Deleted all data 02/05/09 01:15:00 -08:30:00, 02/06/09 03:00:00 – 07:30:00, 02/07/09 04:15:00 – 08:15:00, 02/08/09 05:00:00 – 09:15:00, 02/09/09 06:00:00 – 09:30:00, 02/10/09 07:00:00 – 09:30:00, 02/17/09 00:00:00 – 02:15:00, 02/20/09 03:45:00 – 08:30:00, 02/21/09 04:45:00 – 07:00:00, 02/23/09 05:00:00 – 10:00:00, 02/24/09 05:45:00 – 09:15:00, 02/25/09 07:30:00 – 08:00:00; data suggests sonde was out of the water.

- Suspect turbidity data 02/09/09 09:15:00; value out of range, reason unknown.

SW

- Missing data 02/01/09 00:00:00 – 02/28/09 23:45:00; sonde not deployed.

WT

No data anomalies.

March 1-31, 2009

BC

No data anomalies.

CR

No data anomalies.

DB

- Suspect pH data 03/31/09 10:15:00 – 23:45:00; probe failure, probe subsequently required repair.

- Suspect intermittent negative turbidity data 03/06/09 04:45:00 – 03/31/09 10:00:00; values out of range (See Note #2).

- Suspect turbidity data 03/31/09 12:45:00; value out of range, reason unknown.

HS

- Missing data 03/04/09 14:30:00, associated with sonde exchange.

- Missing data 03/26/09 13:45:00 – 14:45:00, reason unknown.

KB

No data anomalies.

SK

- Deleted all data 03/01/09 22:15:00 – 03/02/09 01:45:00, 03/02/09 22:30:00  - 03/03/09 01:30:00, 03/03/09 23:00:00 – 03/04/09 03:30:00, 03/05/09 01:00:00 – 03:30:00, 03/06/09 03:00:00 – 04:15:00, 03/07/09 03:45:00 – 06:30:00, 03/08/09 05:00:00 – 07:00:00, 03/09/09 06:00:00 – 07:15:00, 03/10/09 07:00:00 – 07:45:00, 03/11/09 20:00:00 – 20:30:00, 03/12/09 20:00:00 – 21:15:00, 03/13/09 20:45:00 – 21:30:00, 03/21/09 03:15:00 – 04:30:00, 03/22/09 03:30:00 – 05:45:00, 03/29/09 22:15:00 – 22:45:00, and 03/30/09 22:15:00 – 23:15:00; data suggests sonde was out of the water.

- Suspect turbidity data 03/05/09 01:30:00, 03/12/09 20:15:00, and 03/21/09 04:00:00; values out of range, reason unknown.

SW

- Missing data 03/01/09 00:00:00 – 03/12/09 11:30:00; sonde not deployed.

WT

No data anomalies.

April 1-30, 2009

BC

No data anomalies.

CR

No data anomalies.

DB

- Suspect turbidity data 04/01/09 12:00:00; 04/04/09 04:15:00; 04/05/09 10:30:00; 04/08/09 00:15:00; 04/09/09 22:30:00; 04/11/09 11:00:00; 04/13/09 21:15:00 – 04/14/09 21:30:00; 04/17/09 07:45:00, 08:00:00, 09:00:00, 09:45:00, 10:15:00, 10:30:00, 11:30:00, 13:00:00 – 14:30:00, 15:00:00, 15:30:00, 15:45:00, 16:15:00 – 16:45:00, 17:15:00, 18:00:00, 18:30:00 – 19:15:00, 21:30:00, 21:45:00, 22:15:00, 22:30:00, 23:00:00, and 23:45:00; 04/18/09 00:00:00, 00:30:00, 01:30:00, 02:45:00, 03:45:00, 04:15:00 – 04:45:00, 05:45:00, 06:00:00, 06:45:00, 07:30:00, 07:45:00, 08:15:00, 08:45:00 – 09:15:00, 10:00:00, 10:45:00 – 11:30:00, 12:00:00, 12:30:00, 13:30:00, 13:45:00, 14:30:00, 15:30:00 – 16:00:00, 16:45:00 – 17:30:00, 18:00:00, 18:30:00, 18:45:00, and 19:30:00 – 04/19/09 06:15:00; and 04/19/09 11:00:00 – 11:30:00; values out of range, turbidity wiper was off of probe at retrieval.

- Suspect larger turbidity values between 04/17/09 08:15:00 – 04/18/09 19:15:00; these values are almost out of range and uncommon, and the turbidity wiper was off of the probe at retrieval.

HS

- Missing data 04/15/09 12:45:00, associated with sonde exchange.

KB

- Suspect turbidity data 04/28/09 22:45:00; value out of range, reason unknown.

SK

- Deleted all data 04/09/09 19:45:00, 04/11/09 20:45:00 – 21:15:00, 04/17/09 00:15:00 – 01:30:00, 04/23/09 17:30:00 – 20:00:00, 04/24/09 18:15:00 – 20:30:00, 04/25/09 18:45:00 – 21:15:00, 04/26/09 19:30:00 – 22:15:00, 04/27/09 20:15:00 – 23:15:00, 04/28/09 21:00:00 – 04/29/09 00:00:00, and 04/29/09 23:00:00 – 23:45:00; data suggests sonde was out of the water.

SW

- Suspect all data 04/14/09 00:00:00 – 04/30/09 23:45:00; sonde guard filled with mud at retrieval.

- Deleted turbidity data 04/14/09 14:15:00 – 14:45:00, 15:15:00, 16:00:00, 16:15:00, and 16:45:00; 04/25/09 16:00:00 – 16:45:00; 04/28/09 19:30:00 – 04/29/09 22:15:00; and 04/30/09 19:15:00 – 23:45:00; sonde guard was filled with mud at retrieval.

WT

No data anomalies.

May 1-31, 2009

BC

No data anomalies.

CR

No data anomalies.

DB

- Missing data 05/11/09 10:45:00; associated with sonde exchange and tube maintenance.

- Shift in D.O. data 05/11/09 11:00:00 associated with sonde exchange (See Note #1).

- Deleted all data 05/11/09 20:45:00 – 22:00:00, 05/12/09 21:45:00 – 22:15:00, 05/19/09 15:30:00 – 05/20/09 04:00:00, 05/24/09 19:15:00 – 20:15:00, and 5/25/09 20:15:00 – 21:30:00; data suggests sonde was out of water.

- Suspect intermittent negative turbidity 05/12/09 02:30:00 – 05/31/09 23:45:00; values out of range (See Note #2).

- Suspect turbidity data 05/07/09 22:45:00; value out of range, reason unknown.

- Suspect negative depth data 05/25/09 19:30:00 – 20:45 and 05/27/09 21:45:00 – 22:45:00; values out of range (See Depth Qualifier).

HS

- Deleted all data 05/26/09 09:00:00; sonde was out of the water, associated with sonde exchange.

KB

- Missing data 05/27/09 13:30:00; associated with sonde exchange, sonde stuck in tube.

- Suspect all data 05/01/09 00:00:00 – 05/27/09 13:15:00; sonde was stuck in tube at retrieval.  All probes and guard covered with barnacles, and wiper brushed shredded at retrieval.

- Deleted all turbidity data 05/08/09 20:45:00 – 05/10/09 13:45:00, 05/15/09 15:15:00 – 05/27/09 13:15:00; probe was covered with barnacles and the wiper brush was shredded at retrieval.

- Deleted D.O. data 05/24/09 10:00:00 – 05/27/09 13:15:00; electrode corroded at retrieval and low post deployment readings (28.5%, 28.0%).

SK

- Deleted all data 05/07/09 17:45:00 – 19:45:00, 05/08/09 18:30:00 – 20:15:00, 05/09/09 18:45:00 – 21:00:00, 05/10/09 19:15:00 – 21:30:00, 05/11/09 20:00:00 – 21:45:00, 05/12/09 20:30:00 – 22:45:00, 05/19/09 02:00:00 – 03:45:00 and 14:30:00 – 22:00:00, 05/23/09 19:00:00 – 19:30:00, 05/24/09 18:45:00 – 21:00:00, 05/25/09 19:30:00 – 21:30:00, 05/26/09 20:30:00 – 22:15:00, and 05/27/09 22:00:00 – 22:30:00; data suggests sonde was out of the water.

- Suspect turbidity data 05/12/09 20:45:00; value out of range, reason unknown.

- Deleted D.O. data 05/18/09 00:00:00 – 05/28/09 11:15:00; low post deployment reading (31.3%, 28.1%) and barnacles covered membrane at retrieval.

- Suspect intermittent negative depth data 05/20/09 00:00:00 – 05/27/09 23:30:00; values out of range (See Depth Qualifier).

SW

- Suspect all data 05/01/09 00:00:00 – 05/11/09 13:00:00; sonde guard filled with mud at retrieval.

- Deleted all turbidity data 05/01/09 00:00:00 – 05/11/09 13:00:00; sonde guard was filled with mud at retrieval.

- Shift in pH 05/11/09 13:15:00 associated with sonde exchange.

WT

No data anomalies.

June 1-30, 2009

BC

No data anomalies.

CR

No data anomalies.

DB

- Deleted all data 06/23/09 20:15:00 – 21:45:00; data suggests sonde was out of the water.

- Suspect D.O. data 06/01/09 00:00:00 – 06/30/09 23:45:00; post deployment readings very low (37.5, 37.0).

- Suspect intermittent negative turbidity data 06/01/09 00:00:00 – 14:30:00; values out of range (See Note #2).

- Suspect turbidity data 06/23/09 21:30:00; value out of range, reason unknown.

- Suspect intermittent negative depth data 06/20/09 17:30:00 – 06/24/09 22:15:00; values out of range (See Depth Qualifier).

HS

No data anomalies.

KB

- Suspect intermittent negative turbidity data 06/17/09 09:15:00 – 06/22/09 21:30:00; values out of range (See Note #2).

- Suspect turbidity data 06/02/09 11:45:00, 06/07/09 02:30:00, 06/16/09 20:30:00, and 06/17/09 08:30:00; values out of range, reason unknown.

SK

- Deleted all data 06/08/09 20:00:00, 06/20/09 17:15:00 – 18:15:00, 06/21/09 17:45:00 – 19:45:00, 06/22/09 19:00:00 – 20:45:00, 06/23/09 19:45:00 – 22:00:00, 06/24/09 21:15:00 – 21:45:00; data suggests sonde was out of the water.

- Missing data 06/17/09 06:15:00 and 08:00:00 – 08:15:00; 06/18/09 05:00:00, 07:00:00, 08:45:00, 09:30:00, 10:00:00, 10:30:00, 11:00:00, 11:45:00, 13:00:00, 14:30:00 – 14:45:00, 16:00:00 – 16:15:00, 16:45:00 – 17:45:00, 18:15:00, 18:45:00, 19:15:00, 19:45:00 – 20:45:00, and 21:15:00 – 06/19/09 02:15:00; 06/19/09 02:45:00 – 04:00:00, 04:30:00 – 08:00:00, and 08:30:00 – 10:45:00; battery failure.

- Suspect D.O. data 06/19/09 11:00:00 – 06/30/09 23:45:00; post deployment values low (64.8, 64.6).  Barnacles were all over the probe, but not the membrane; the wiper apparatus protected the membrane.

- Suspect intermittent negative depth data 06/19/09 16:00:00 – 06/25/09 22:15:00; values out of range (See Depth Qualifier).

- Suspect turbidity data 06/06/09 18:30:00, 06/09/09 21:30:00, 06/15/09 22:30:00, 06/17/09 21:30:00; values out of range, reason unknown.

SW

- Suspect turbidity data 06/30/09 07:45:00; value out of range, reason unknown.

WT

No data anomalies.

July 1-31, 2009

BC

No data anomalies.

CR

- Missing data 07/15/09 10:15:00; associated with sonde exchange.

DB

- Suspect D.O. data 07/1/09 00:00:00 – 07/14/09 12:00:00, low post deployment values (37.5, 37.0).

- Suspect intermittent negative turbidity data 07/14/08 12:15:00 – 07/30/09 11:30:00; values out of range (See Note #2).

- Shift in D.O., depth, and pH data 07/14/09 12:15:00 associated with sonde exchange (See Note #1).

HS

No data anomalies.

KB

- Suspect D.O. data 07/06/09 04:00:00 – 07/15/09 10:45:00; mud found on membrane at retrieval and low post deployment readings (70.1, 70.5).

- Shift in D.O. 07/15/09 11:00:00 associated with sonde exchange (See Note #1).

SK

- Suspect turbidity data 07/17/09 08:45:00; value out of range, reason unknown.

- Deleted all data 07/19/09 16:45:00 -18:45:00, 07/20/09 17:30:00 – 20:00:00, 07/21/09 18:00:00 – 20:45:00, 07/22/09 19:00:00 – 21:30:00, and 07/23/09 20:00:00 – 21:30:00; data suggests sonde was out of the water.

- Suspect D.O. data 07/01/09 00:00:00 – 07/17/09 09:30:00; post deployment values low (64.8, 64.6).  Barnacles were all over the probe, but not the membrane; the wiper apparatus protected the membrane.

- Shift in D.O. data 07/17/09 09:45:00 associated with sonde exchange (See Note #1).

SW

- Suspect turbidity data 07/01/09 00:00:00 – 07/14/09 09:45:00; sonde guard filled with mud at retrieval.

- Shift in D.O. data 07/30/09 09:45:00 associated with sonde exchange (See Note #1).

WT

- Missing data 07/29/09 09:45:00; associated with sonde exchange, connection issues.

August 1-31, 2009

BC

No data anomalies.

CR

No data anomalies.

DB

- Shift in pH 08/28/09 10:00:00 associated with sonde exchange.

- Suspect intermittent turbidity data 08/22/09 03:00:00 – 08/28/09 04:45:00; values out of range, reason unknown.

HS

No data anomalies.

KB

No data anomalies.

SK

- Missing data 08/05/09 08:45:00 – 08/31/09 23:45:00; sonde site hit by lightning and/or vandalized.

SW

- Deleted turbidity data 08/31/09 09:45:00 – 08/31/09 23:45:00; probe failure.

WT

No data anomalies.

September 1-30, 2009

BC

No data anomalies.

CR

No data anomalies.

DB

- Suspect turbidity data 09/13/09 01:00:00 and 01:30:00; 09/16/09 03:15:00, 03:30:00, and 05:15:00; 09/17/09 22:15:00; 09/18/09 02:45:00 and 03:30:00; 09/19/09 00:15:00; 09/20/09 00:00:00; 09/21/09 20:45:00 and 21:45:00; 09/22/09 00:00:00; 09/23/09 03:30:00 and 06:00:00; 09/24/09 04:15:00 – 04:45:00; 09/25/09 03:15:00, 03:45:00, 14:45:00, and 18:45:00; 09/26/09 04:00:00, 05:30:00, 21:30:00, 22:30:00, and 23:00:00; 09/27/09 19:30:00, 22:45:00, 23:15:00; 09/28/09 04:45:00, 19:30:00, and 23:30:00; 09/29/09 06:30:00, 07:15:00, and 19:00:00; and 09/30/09 04:00:00, 06:00:00, and 17:30:00; values out of range, reason unknown.

HS

No data anomalies.

KB

- Suspect D.O. data 09/17/09 12:45:00 – 09/30/09 13:00:00; post deployment readings low (79.1, 76.4).

SK

- Missing data 09/01/09 00:00:00 – 09/08/09 09:15:00; sonde site hit by lightning and/or vandalized.

- Suspect D.O. data 09/08/09 09:30:00 – 09/30/09 23:45:00; possible calibration offset (post deployment readings 158.1, 159.3) and electrodes corroded at retrieval.

- Deleted all intermittent negative turbidity data between 09/12/09 13:15:00 and 09/30/09 23:45:00; result of a calibration offset (post deployment value of -4.9 NTU in 0.0 NTU water).

SW

- Suspect D.O. data 09/08/09 15:00:00 – 09/09/09 08:45:00; values out of range, thought to be due to the presence of hydrogen sulfide.

- Missing data 09/29/09 14:00:00, 14:15:00, 15:45:00, and 22:00:00; 09/30/09 01:00:00 – 02:00:00, 03:00:00 – 03:30:00, 04:30:00 – 04:45:00, 07:15:00 – 07:30:00, 10:00:00, 13:00:00 – 14:45:00, and 15:30:00 – 23:45:00; battery failure.

- Deleted turbidity data 09/01/09 00:00:00 – 09/30/09 23:45:00; probe failure.

WT

- Missing data 09/10/09 08:30:00 and 08:45:00; associated with sonde exchange, connection issues.

October 1-31, 2009

BC

No data anomalies.

CR

No data anomalies.

DB

- Missing data 10/05/09 10:15:00; tube replacement/maintenance.

- Suspect turbidity data 10/01/09 05:00:00, 10:15:00, 21:30:00, and 22:15:00; 10/02/09 05:45:00, 06:00:00, and 22:30:00; 10/03/09 00:45:00, 04:00:00, 06:00:00, 22:00:00, 22:45:00; values out of range, reason unknown.

- Suspect intermittent negative turbidity data 10/28/09 11:30:00 – 10/31/09 23:45:00; values out of range (See Note #2).

- Suspect D.O. data 10/05/09 10:30:00 – 10/28/09 10:30:00; low post deployment reading.

- Shift in D.O. data 10/28/09 10:45:00 associated with sonde exchange (See Note #1).

HS

No data anomalies.

KB

- Suspect intermittent negative turbidity data 10/22/09 09:15:00 – 10/31/09 07:30:00; values out of range (See Note #2).

SK

- Deleted all pH data 10/08/09 09:15:00 – 10/29/09 09:15:00; no probe installed for this deployment.

- Deleted all turbidity data 10/17/09 20:30:00 – 10/29/09 09:15:00; wiper off of probe at retrieval and guard full of mud.

- Deleted all turbidity data 10/29/09 09:15:00 – 10/31/09 23:45:00; negative data attributed to a calibration offset, post deployment value was -16.0 NTU in 0.0 NTU water.

- Deleted intermittent negative turbidity data 10/01/09 00:00:00 – 10/07/09 00:00:00; values out of range, possibly due to calibration offset.  Post deployment reading -4.9 NTU in 0.0 NTU water.

SW

- Missing data 10/01/09 00:00:00 – 02:30:00, 03:00:00 – 06:45:00, 07:15:00 – 11:30:00, and 12:00:00; battery failure.

- Suspect D.O. data 10/12/09 12:45:00 – 10/28/09 12:30:00; post deployment readings very low (35.4, 35.6).

- Suspect turbidity data 10/23/09 14:00:00, 10/25/09 08:00:00, and 10/28/09 10:00:00; values out of range, reason unknown.

WT

No data anomalies.

November 1-30, 2009

BC

No data anomalies.

CR

No data anomalies.

DB

- Suspect intermittent negative turbidity data 11/01/09 00:00:00 – 11/13/09 12:30:00; values out of range, reason unknown (See Note #2).

- Suspect D.O. values 11/13/09 12:45:00 – 11/30/09 23:45:00; post deployment values low (57.3, 57.1).

- Shift in pH data 11/13/09 12:45:00 associated with sonde exchange (See Note #1).

HS

No data anomalies.

KB

No data anomalies.

SK

- Deleted all data 11/05/09 10:15:00 and 10:30:00 and 11/06/09 08:45:00 – 11:30:00; values suggest sonde was out of the water.

- Shift in D.O. and pH data 11/16/09 13:15:00 associated with sonde exchange (See Note #1).

- Deleted all turbidity data 11/01/09 00:00:00 – 11/16/09 13:00:00; negative data attributed to a calibration offset, post deployment value was -16.0 NTU in 0.0 NTU water.

- Suspect turbidity data 11/26/09 07:45:00; value out of range, reason unknown.

SW

- Deleted D.O. data 11/13/09 10:45:00 – 11/30/09 23:45:00; hole in membrane at retrieval and post deployment readings low (29.3, 29.1).

WT

No data anomalies.

December 1-31, 2009

BC

No data anomalies.

CR

No data anomalies.

DB

- Suspect D.O. values 12/1/09 00:00:00 – 12/11/09 12:45:00; low post deployment values (57.3, 57.1).

- Shift in D.O. and pH data 12/11/09 13:00:00 associated with sonde exchange (See Note #1).

- Deleted intermittent negative turbidity data 12/11/09 13:00:00 – 12/31/09 23:45:00; probe failure, wiper not parked at 180o and not responding at retrieval.

HS

No data anomalies.

KB

No data anomalies.

SK

- Deleted all data 12/01/09 07:00:00 and 07:15:00; 12/04/09 08:15:00 – 10:30:00; 12/05/09 09:00:00 – 11:00:00; 12/06/09 09:30:00 – 11:30:00; 12/14/09 06:15:00 and 06:30:00; 12/15/09 06:30:00 – 07:30:00; 12/16/09 06:00:00 – 09:45:00; 12/17/09 06:00:00 – 10:45:00; 12/19/09 08:45:00 – 10:30:00; 12/20/09 08:45:00 – 11:00:00; 12/21/09 09:30:00 – 10:00:00; 12/29/09 03:30:00 – 09:00:00; 12/30/09 04:30:00 – 08:30:00; and 12/31/09 06:00:00 – 08:45:00; data suggests sonde was out of the water.

- Shift in pH data 12/15/09 12:45:00 associated with sonde exchange (See Note #1).

- Suspect intermittent negative turbidity data 12/15/09 13:15:00 – 12/16/09 10:15:00; values out of range (See Note #2).

- Suspect intermittent negative depth data 12/18/09 08:15:00 – 12/19/09 11:15:00; values out of range (See Depth Qualifier).

SW

- Deleted D.O. data 12/01/09 00:00:00 – 12/11/09 10:15:00; hole in membrane at retrieval and low post deployment values (29.9, 29.1).

- Deleted turbidity data 12/26/09 12:00:00 – 12/31/09 23:45:00; values out of range, mud at bottom of sonde guard at retrieval.

WT

No data anomalies.

12) Post deployment information:

End of deployment Post-calibration readings in standard solutions prior to probe cleaning.

Table 3.  Post-deployment readings of 6600 EDS sondes deployed at the Seahorse Key site during 2009.

pH	        Temp (C)       SC (mS/cm)        DO %	        Turbidity (NTU)     Depth (m)

Date/Std. 	7.00		n/a		50.00		100.0		    0.0		            n/a

01/26/09	6.64		N/A		49.61		99.4, 99.4	    0.0			0.088

03/05/09	7.00		N/A		50.65		85.9, 86.5	    0.5			0.160

04/01/09	7.02		N/A		49.11		44.9, 45.0	    0.2			0.018

04/28/09	7.00		N/A		46.98		96.0, 96.8	    0.2			0.017

05/28/09	7.29		N/A		45.21		31.3, 28.1	    0.4			0.001

06/19/09	6.85		N/A		50.35		124.9, N/A	    0.4			0.014

07/17/09	6.90		N/A		43.12		64.6, 64.8	    0.4			0.040

08/05/09	6.88		N/A		49.35		114.3, 114.3	    0.5			0.069

08/25/09	N/A		N/A		N/A		N/A, N/A	    N/A			N/A

10/08/09	6.74		N/A		50.97		158.1, 159.3	   -4.9			0.041

10/29/09	7.40		N/A		51.16		82.9, 83.2	    0.8			0.036

11/16/09	7.02		N/A		49.52		70.4, 70.7	   -16.0			0.045

12/15/09	8.34		N/A		49.74		75.9, 70.9	    0.3			-0.105

01/25/10	7.06		N/A		50.64		92.7, 92.7	    1.1			-0.017

Table 4.  Post-deployment readings of all 6600 EDS sondes deployed at the King’s Bay site during 2009.

pH	       Temp (C)	   SC (mS/cm)	      DO % sat.	        Turbidity (NTU)     Depth (m)

Date/Std.	7.00		n/a		50.00		100.0		    0.0		         n/a

01/22/09	7.21		N/A		50.05		104.8, 103.0	    0.5			0.121

02/18/09	7.06		N/A		49.50		91.5, 91.3	    1.3			0.062

03/11/09	7.02		N/A		50.40		99.5, 99.7	    0.3			0.081

03/26/09	7.36		N/A		50.03		100.8, 100.7	    0.0			0.058

04/16/09	7.06		N/A		49.78		99.2, 99.2	    0.0			0.057

05/27/09	7.28		N/A		46.20		28.5, 28.0	    2.9			-0.008

06/17/09	7.08		N/A		49.80		99.5, 99.6	    -0.9			0.046

07/15/09	7.15		N/A		50.01		70.1, 70.5	    -1.8			0.079

08/04/09	7.03		N/A		49.82		100.3, 101.3	    0.4			0.073

08/18/09	7.07		N/A		50.45		142.3, 140.3	   -0.4			0.051

09/17/09	7.06		N/A		49.66		90.3, 90.3	    0.7			0.041

09/30/09	7.07		N/A		50.25		79.1, 76.4	    0.3			0.021

10/22/09	7.00		N/A		49.14		98.6, 98.7	    0.6			0.028

11/12/09	7.03		N/A		49.08		96.7, 96.5	    -1.0			0.052

12/10/09	7.06		N/A		49.71		97.4, 96.5	    0.7			-0.109

01/22/10	N/A		N/A		N/A		N/A, N/A	    N/A			N/A

Table 5.  Post-deployment readings of 6600 EDS sondes deployed at the Dekle Beach site during 2009.

pH	        Temp (C)       SC (mS/cm)        DO %	        Turbidity (NTU)     Depth (m)

Date/Std. 	7.00		n/a		50.00		100.0		    0.0		            n/a

01/08/09	7.01		N/A		49.93		92.8, 92.1	    1.0			0.102

02/26/09	7.27		N/A		50.59		92.8, 92.1	    0.1			0.172

03/31/09	7.03		N/A		49.92		134.7, 134.6	   -0.4			0.021

05/11/09	7.78		N/A		49.20		67.1, 67.0	    0.8			0.072

06/01/09	7.07		N/A		51.10		119.0, 118.7	   -0.1			0.035

07/14/09	8.02		N/A		50.43		37.5, 37.0	    0.2			0.080

07/30/09	6.93		N/A		51.00		91.5, 91.6	   -0.5			0.090

08/28/09	7.88		N/A		49.95		87.8, 87.9	    0.4			0.038

10/05/09	7.03		N/A		49.83		105.7, 107.0	    3.6			0.018

10/28/09	7.45		N/A		49.94		53.1, 53.2	    0.3			0.026

11/13/09	6.97		N/A		49.24		102.8, 102.9	    -2.2			0.044

12/11/09	8.20		N/A		50.08		57.3, 57.1	    0.5			-0.109

01/12/10	7.00		N/A		49.88		-50.0, -50.0	   -6.2			-0.011

Table 6.  Post-deployment readings of 6600 EDS sondes deployed at the Suwannee River site during 2009.

pH	        Temp (C)       SC (mS/cm)        DO %	        Turbidity (NTU)     Depth (m)

Date/Std. 	7.00		n/a		50.00		100.0		    0.0		            n/a

03/31/09	6.98		N/A		50.28		95.6, 95.9	    0.3			0.006

05/11/09	7.73		N/A		50.96		67.1, 67.1	    0.0			0.070

06/01/09	6.93		N/A		50.09		104.1, 103.8	    0.2			0.023

07/14/09	7.02		N/A		49.53		58.1, 58.2	    0.4			0.091

07/30/09	6.90		N/A		50.05		100.1, 100.3	    1.0			0.094

08/31/09	6.89		N/A		50.11		65.0, 65.0	    0.2			0.047

10/01/09	7.06		N/A		45.29		108.7, 109.0	    4.1			0.016

10/28/09	6.97		N/A		50.99		35.4, 35.6	    0.0			0.036

11/13/09	7.08		N/A		53.49		92.4, 92.4	   -0.4			0.074

12/11/09	6.89		N/A		49.85		29.3, 29.1	    0.3			-0.105

01/12/10	7.01		N/A		49.86		115.7, 116.3	    0.9			-0.008

13) Other remarks/notes:

- Calibration of DO was performed on the day of deployment.  Two DO values are recorded during the post-deployment evaluation process.

- This style of metadata was formerly used by National Estuarine Research Reserve program; more specifically, this report was modeled after metadata reports created at the Guana Tolomato Matanzas National Estuarine Research Reserve.

- Copies of calibration/deployment logs can be obtained through the Principal Investigator.